|  | FORM ACB-PC1 ACB WITNESS TEST PROCEDURE AND CHECKLIST AHRI CERTIFICATION PROGRAM FOR ACTIVE CHILLED BEAMS |
| --- | --- |

ACB Participant:		____________________________________________________

Date of Test: 		____________________________________________________

Witness Test Number: 		____________________________________________________

Test Results (P/F):

Instructions to Laboratory Representatives: The following Witness Test Procedure and Checklist is to be used by the Representative Contracted by AHRI (Representative) to document all steps required to conduct witness tests for the Air-Cooled (ACB) Water Chilling Packages Certification Program. This Witness Test Procedure and Checklist is to be used in conjunction with the AHRI General Operations Manual (OM) for AHRI Certification Programs, the ACB OM and the latest edition of AHRI Standard 1240 (I-P) or AHRI Standard 1241 (SI) and ASHRAE 200 . Where the AHRI General OM, the ACB OM and this Witness Test Procedure and Checklist differ, the other documents shall prevail. The completed checklist shall be submitted to AHRI upon the completion of the witness test.

If the Participant’s test facility (Facility) fails to comply with the conditions listed below:

- Notify AHRI

- Notify Participant

| Representative initials | data point (if applicable) |  |
| --- | --- | --- |

| Pre-Visit |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | When setting up for a trip to Witness Test an Active Chilled Beam for AHRI Certification Programs, several information items are required. Verify the following information as received from AHRI: - Letter of Authorization (AHRI Selection Letter, refer to Section 3 of the ACB OM) from AHRI for the Third-Party Laboratory contracted by AHRI (Laboratory) to witness the Certification Test - Copy of the latest AHRI Facility Inspection Approval Report - Test Site Location - Participant’s Contact(s) with Phone and Fax Numbers and email - Active Chilled Beam Model Number - Model Performance Ratings |  |  |  |  |
|  |  |  | Verify the following information as received from the Participant: - Verification of Power Input instrumentation requirements - Mutually agreed upon test date of when to arrive at the Participant’s specified facility - Directions to the Participant’s specified facility - Estimate of when testing will be completed |  |  |  |  |
|  |  |  | Inform the Participant that the unit can be set-up as per the latest edition of the ACB OM when the Representative arrives |  |  |  |  |
|  |  |  | Ensure that all Laboratory instrumentation has current calibration stickers. Conduct a quick test to verify that the maximum deviation of the temperature probes from the highest to the lowest readings is within 0.10°F [0.055°C]. This can be done by placing the probes through a wooden board or 1-inch rubber insulation and setting it on a bucket of near ambient water. Stir the water while checking the temperature readings from each probe. Document the temperature probe readings and recalibrate if necessary. Extreme care must be exercised when handling the temperature probes. Leave the rubber protective sleeves on the probes until you are ready to insert them into the thermowells or into the water streams. |  |  |  |  |
| Minimum Laboratory instrumentation required: |  |  |  |  |  |  |  |
|  |  |  | Water temperature measurement device |  |  |  |  |
|  |  |  | Air temperature measurement device |  |  |  |  |
|  |  |  | Barometer |  |  |  |  |
| Representative initials |  | data point (if applicable) |  |  |  |  |  |

| Day of Test |  |  |
| --- | --- | --- |
|  |  | Upon arrival, request a brief meeting of the involved parties to review the following information: - Safety rules at the test site - Confirm test conditions - Test set-up; location of Laboratory instrumentation |
|  |  | Representative shall unpack and visually inspect all Laboratory instrumentation |
|  |  | Representative shall install all Laboratory instrumentation |
|  |  | Representative shall visually inspect and confirm the test sample active chilled beam nameplate against the AHRI Selection Letter to confirm identity of the active chilled beam under test. |
|  |  | For water pressure drop, the “Zero” measurements shall be recorded and shall be zero, within a tolerance of ± 1.0% of the full scale value of the calibration range of the measurement system. This verification can be completed before or after the test has been conducted. |
|  |  | For water flow, the “Zero” measurements shall be recorded and shall be zero, within a tolerance of ± 1.0% of the full scale value of the calibration range of the measurement system. This verification can be completed before or after the test has been conducted. |
|  |  | Representative shall verify the type of data reported: 1. Five (5) minute average point (continuous) or 2. Instantaneous data point |
|  |  | Representative shall authorize starting the test sample active chilled beam |

| Representative Initials | data point (if applicable) |  |
| --- | --- | --- |

| Test |  |  |  |  |
| --- | --- | --- | --- | --- |
| Representative shall confirm and record that the Laboratory instrumentation is in agreement with the Facility’s instruments. If these criteria are not met, interchange the Laboratory instrumentation to confirm they are correct. Should any of these criteria not be met, steps should be taken to ensure that the discrepancy is resolved. In cases where the discrepancy cannot be resolved and no damage is apparent, the Laboratory instrumentation data for that parameter shall be used. |  |  |  |  |
|  |  | - The water temperature difference across the ACB as determined by the Facility’s instrumentation is within 0.2ºF from the Laboratory instrumentation |  |  |
|  |  | - The Facility and Laboratory temperatures values for entering water temperature are no more than 0.2ºF different |  |  |
|  |  | - The Water Pressure Drop reading determined by the Facility’s instrumentation is within 2% of the Laboratory’s values |  |  |
|  |  | - The Facility and Laboratory barometric pressure values differ by no more than 0.02 psi |  |  |
|  |  | The facility operator is to inform the Representative when they are ready to initiate the test so the Representative can record test data simultaneously. The Representative shall verify that the Facility test conditions remain within the allowable tolerances for the duration of the 15 minute test period, by checking the Representative’s test condition measurements at intervals not to exceed 15 seconds. |  |  |
| This checklist & all supporting documents have been reviewed & approved by: ________________________________________ PRINTED NAME OF REPRESENTATIVE ________________________________________ SIGNATURE OF REPRESENTATIVE ________________________________________ DATE |  |  |  |  |