## Organization
LABORATORY REVIEW AND APPROVAL
SUMMARY FORM
Drawings and Photos
Capacity
| Figure | 12 | 15
Pitot Traverse
Nozzles
Barometer
Electric Meters
Pressure
Speed
Temperature
Torque
Check Tests and Results
Leakage
Personnel
These checklists are used by Facility staff to ensure that the requirements of AMCA Standard 210 are documented during the lab accreditation process. They reflect the calibration practices of the Facility Laboratory and may have to be modified for different instrumentation. Please refer to AMCA Standard 210 for more complete information.

## Capacity
LABORATORY REVIEW AND APPROVAL
CHAMBER CAPACITY
Cross-sectional area of chamber: |  |  | Units:
Check the appropriate box(es) below. All yellow boxes will be automatically calculated.
| Figure 12 from AMCA Standard 210 |  |  | FALSE
Dimension J:
Axis of rotation perpendicular to flow:
Cross-sectional area / 9:
Outlet area based on a diameter equal to dimension J:
Maximum fan outlet area:
Axis of rotation parallel to flow:
Cross-sectional area / 16:
Outlet area based on a diameter equal to dimension J / 2:
Maximum fan outlet area:
| Figure 15 from AMCA Standard 210 |  |  | FALSE
Cross-sectional area / 5, or maximum fan inlet area:

## Figure 12
LABORATORY REVIEW AND APPROVAL
FIGURE 12 FROM AMCA 210
Please provide drawings to document the following:
Static Pressure Taps
Dimensions conform to Figure 2A
Piezometer Rings
Plane 5
1.25 to 1.75 inches from the nozzle wall
At least four pressure taps
Plane 6
1.25 to 1.75 inches from the nozzle wall
At least four pressure taps
Plane 7
0.5J from the chamber entrance
At least four pressure taps
Nozzle Wall
The centerlines of the nozzles are at least 1.5 nozzle exit diameters from the chamber
wall.
The minimum distance between the centers of any two nozzles in simultaneous use are
3 times the exit diameter of the larger nozzle.
The nozzle wall is at least 0.5M from the upstream settling means.
Settling Means
The downstream settling means are at least 2.5 exit diameters from the exit face of
the largest nozzle.
Velocity Traverse
Conduct velocity traverses to determine if the settling means are sufficient. Conduct one
traverse near Plane 5 (0.1M downstream of the settling means) with the auxiliary fan at its
maximum and an opening in the chamber slightly larger than that of a typical fan with that air-
flow rate.
Conduct another traverse near Plane 7 with the same airflow rate and chamber opening to
check for reverse flow.
Conduct another traverse near Plane 6 to check for reverse flow.
Do not purposely select a flow rate that will give a maximum velocity under 400 feet per minute.
Please attach:
Traverse results
Chamber opening dimensions
Nozzle combination
Please demonstrate that:
The maximum velocity at 0.1M downstream from the settling means does not exceed
the average velocity by more than 25% when the maximum velocity is greater than 400
feet per minute.
The maximum reverse jet velocity at Plane 7 does not exceed 10% of the average
velocity at Plane 2.
The maximum reverse jet velocity at Plane 6 does not exceed 10% of the average
nozzle velocity.
Recorded Data
Please indicate the make, model number and serial number of the instrument(s) used to
record the data shown below. Indicate any exceptions allowed by AMCA 210.
td0
tw0
pb
N
T
W
DP
td5
PS5
PS7
Example:
T |  | RTD: Omega Model No., Serial No. - RTD Meter: Newport Electronics Model No., Serial No. DAS: National Instruments Model No., Serial No. - PC: HP Model No., Serial No.
RTD
RTD Meter
Data Acquisition System (DAS)
PC

## Figure 15
LABORATORY REVIEW AND APPROVAL
FIGURE 15 FROM AMCA 210
Please provide drawings to document the following:
Static Pressure Taps
Dimensions conform to Figure 2A
Piezometer Rings
Plane 5
1.25 to 1.75 inches from the nozzle wall
At least four pressure taps
Plane 6
1.25 to 1.75 inches from the nozzle wall
At least four pressure taps
Total Pressure Tube
Dimensions conform to Figure 2B
The minimum distance from the chamber wall, floor and ceiling is 0.1M.
Plane 8
The minimum distance from the settling means is 0.2M.
The minimum distance from the chamber exit is 0.3M.
Nozzle Wall
The centerlines of the nozzles are at least 1.5 nozzle exit diameters from the chamber
wall.
The minimum distance between the centers of any two nozzles in simultaneous use are
3 times the exit diameter of the larger nozzle.
The nozzle wall is at least 0.5M from the upstream settling means.
Settling Means
The downstream settling means are at least 2.5 exit diameters from the exit face of
the largest nozzle.
The downstream settling means are at least 0.5M from the chamber exit.
The upstream settling means are at least 0.2M from the variable supply system.
Velocity Traverse
Conduct velocity traverses to determine if the settling means are sufficient. Conduct one
traverse near Plane 5 (0.1M downstream of the settling means) with the auxiliary fan at its
maximum.
Conduct another traverse near Plane 8 (0.1M downstream of the settling means) with a typical
nozzle combination in the mid-range of the chamber capacity with a typical ΔP.
Conduct another traverse near Plane 6 to check for reverse flow with the above mid-range
nozzle combination.
Do not purposely select a flow rate that will give a maximum velocity under 400 feet per minute.
Please attach:
Traverse results
Chamber opening dimensions
Nozzle combination
Please demonstrate that:
The maximum velocity at 0.1M downstream from the settling means at Plane 5 does
not exceed the average velocity by more than 25% when the maximum velocity is
greater than 400 feet per minute.
The maximum velocity at 0.1M downstream from the settling means at Plane 8 does
not exceed the average velocity by more than 25% when the maximum velocity is
greater than 400 feet per minute.
The maximum reverse jet velocity at Plane 6 does not exceed 10% of the average
nozzle velocity.
Recorded Data
Please indicate the make, model number and serial number of the instrument used to record
the data shown below. Indicate any exceptions allowed by AMCA 210.
td0
tw0
pb
N
T
W
DP
td5
PS5
td8
Pt8
Example:
T |  | RTD: Omega Model No., Serial No. - RTD Meter: Newport Electronics Model No., Serial No. DAS: National Instruments Model No., Serial No. - PC: HP Model No., Serial No.
RTD
RTD Meter
Data Acquisition System (DAS)
PC

## Nozzles
LABORATORY REVIEW AND APPROVAL
NOZZLES
Nozzle Identification
Throat Entrance
| Position | Diameter, D | D < 1.002 Mean D? |  | D > 0.998 Mean D?
0°
45°
90°
135°
Mean D:
1.002 Mean D:
0.998 Mean D:
Nozzle Exit
| Position | Diameter, D | D < 1.002 Mean D? |  | D > 0.998 Mean D?
0°
45°
90°
135°
Mean D:
1.002 Mean D:
0.998 Mean D:
| Exit D < Entrance D: |  |  | Entrance D < 1.002 Exit D:
Nozzle measurements OK?
Nozzle was checked to be smooth, square, sharp edged, and free from nicks.
Please provide a description of the instrument used to measure the nozzles.
Please provide a description as to how the angle between diameter measurements was
verified.
Please confirm that the nozzles were measured in situ.
Please confirm that actual average diameters are used to calculate flow rate as opposed to
nominal diameter.
Please submit a copy of this form for each chamber nozzle.

## Barometer
LABORATORY REVIEW AND APPROVAL
BAROMETER
| Reference Standard |  | Instrument Under Calibration
Manufacturer:
Model:
Serial No.:
Range:
Graduation:
Remarks:
Reading: |  | in Hg |  | in Hg
Difference: |  | in Hg
OK?
This form must be modified when calibrating indiviual components of a data acquisition
system and combining the uncertainties.

## Electric Meters
LABORATORY REVIEW AND APPROVAL
ELECTRIC METERS
| Reference Standard | Instrument Under Calibration
Manufacturer:
Model:
Serial No.:
Range:
Graduation:
Remarks:
Please complete this form and attach calibration records for each electric meter.
Most power meters come with a display and the calibration will be of a known input against
the display. If the power meter will be used with a data acquisition system the display must
match the value from the DAS.
This form must be modified when calibrating individual components of a data acquisition
system and combining uncertainties.
Please provide documentation on how you plan to maintain supply voltage to within 1% of the
motor calibration voltage, if your Facility intends to use the Motor Calibration Method.
Please note that the serial number of the calibrated motor must match the serial number of the
fan motor.
Calibration of the wattmeter must include any clamp on amp transducer used with the watt-
meter.
Calibration of wattmeters must include power factor. A factor between 0.8 and 0.9 is
recommended.

## Pressure
LABORATORY REVIEW AND APPROVAL
AMCA STANDARD 210
PRESSURE
|  | Reference Standard |  |  | Instrument Under Calibration
Manufacturer:
Model:
Serial No.:
Range:
Graduation:
Remarks:
| Reference |  | Instrument Under |  |  | Percent
Point | Standard |  | Calibration | Difference |  | Difference | OK?
1
2
3
4
5
6
7
8
9
10
11
The reference standard must have a traceable calibration and a minimum accuracy of a typical
micromanometer (0.004 in. w.g.).
This form is formatted for end-to-end calibration. It must be modified when calibrating
individual components of a data acquisition system and combining the uncertainties.
This form must be completed for each pressure measuring device.
Note on C26: Opalka, Mary: To be filled out by AMCA?

## Speed
LABORATORY ACCREDITATION
AMCA STANDARD 210
SPEED
|  | Reference Standard |  |  | Instrument Under Calibration
Manufacturer:
Model:
Serial No.:
Range:
Graduation:
Remarks:
| Reference |  | Instrument Under |  |  | Percent
Point | Standard |  | Calibration | Difference |  | Difference | OK?
1
2
3
4
5
6
7
8
9
10
11
This form is formatted for end-to-end calibration. It must be modified when calibrating
individual components of a data acquisition system and combining the uncertainties.
This form must be completed for each speed measuring device.

## Temperature
LABORATORY ACCREDITATION
AMCA STANDARD 210
TEMPERATURE
|  | Reference Standard |  |  |  |  | Instrument Under Calibration
Manufacturer:
Model:
Serial No.:
Range:
Graduation:
Remarks:
|  | Reference |  | Instrument Under
|  | Standard |  | Calibration |  | Difference |  |  | OK?
|  |  | Velocity over wet bulb: |  |  |  |  | ft/min
OK?
Please attach:
A drawing showing dry bulb mounted upstream of wet bulb
This form is formatted for end-to-end calibration. It must be modified when calibrating
individual components of a data acquisition system and combining the uncertainties.
This form must be completed for each temperature measuring device.

## Torque
LABORATORY ACCREDITATION
AMCA STANDARD 210
TORQUE
|  | Reference Standard |  |  | Instrument Under Calibration
Manufacturer:
Model:
Serial No.:
Range:
Graduation:
Remarks:
|  | Clockwise rotation |  |  | Counterclockwise rotation
| Applied |  | Instrument Under |  |  | Percent
Point | Torque |  | Calibration | Difference |  | Difference | OK?
1
2
3
4
5
6
7
8
9
10
11
Each lab must have the ability to calibrate the torque measurement system prior to each test.
Please attach:
A drawing of the torque arm
Calibration records of the calibration weights
This form is formatted for end-to-end calibration.
Please submit a copy of this form for each torque measurement device calibrated in each
rotational direction.

## Test Results
LABORATORY ACCREDITATION
AMCA STANDARD 210
TEST RESULTS
The list below was developed from AMCA Standard 210 and Publication 111.
Standard 210
Not applicable
O.K. | Missing |  | Test Results
Calculated Results
Raw test data
Test fan description, including appurtenances
Test setup - test figure and installation type
List of test instruments
Name and location of laboratory
Names of test personnel
Performance Curve
Flow rate as abscissa on curve
Pressure (static, total, or both) and power as ordinates on curve
Speed listed if corrected to a constant, or else plotted on curve
Fan air density listed if corrected to a constant, or else plotted on curve
Barometric pressure listed on curve if fan pressures exceed 10 in. wg
Test points identified on curve (series of circled points)